## Slide 1: Ozone Calibration Data

## Slide 2: Changes in Data
The monitor is now reading lower concentrations during the initial peak of a run.
  However, rate of increase of ozone concentration seems to be increasing. (It takes longer to stabilize.)
AirU sensors are showing less resolution in peaks than previously. This does not necessarily seem to be linked to low concentrations of ozone.
In some cases, the NOX value is increasing rather than decreasing.

## Slide 3: O3_2018-06-04-14-32_C (A run with casing that gave us ‘normal’ data.)
87 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 88.9 ppb | 96.7 ppb | 101.3 ppb | 
S5 NOX Value | 521 | 470 | 414 | 99.6%
S8 NOX Value | 1113 | 980 | 879 | 98.9%
S12 NOX Value | 1064 | 995 | 912 | 99.0%
Chart: 
Sensor 12 NOX values

## Slide 4: O3_2018-05-23-12-57_nofan (A run without the fan that gave us ‘normal’ data).
87 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 127.2 ppb | 148.5 ppb | 157.7 ppb | 
S5 NOX Value |  |  |  | 91.1%
S8 NOX Value |  |  |  | 96.0%
S12 NOX Value | 721 | 530 | 392 | 95.5%
Chart: 
Sensor 12 NOX values

## Slide 5: O3_2018-06-08-11-00_C1000ppb
 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 37 ppb | 60 ppb | 86 ppb | 
S5 NOX Value | 1070 | 950 | 740 | 84.0%
S8 NOX Value | 1640 | 1712 | 1434 | 77.8%
S12 NOX Value | 1525 | 1455 | 1253 | 81.7 %
O3_2018-06-08-14-36_nofan1000ppb
 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 210 ppb | 246 ppb | 275 ppb | 
S5 NOX Value | 139 | 12 | 11 | 89.6%
S8 NOX Value | 159 | 129 | 79 | 96.7%
S12 NOX Value | 394 | 203 | 112 | 98.1%
Chart: 
Chart: 
Sensor 12 NOX values
Sensor 12 NOX values

## Slide 6: O3_2018-06-11-10-10_Cnofan
 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 34.4 ppb | 53 ppb | 76 ppb | 
S5 NOX Value | 1025 | 1000 | 920 | 95.3%
S8 NOX Value | 1655 | 1823 | 1843 | 8.9%
S12 NOX Value | 1431 | 1581 | 1588 | 3.1%
O3_2018-06-11-13-20_nofan
 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 106 ppb | 145 ppb | 171 ppb | 
S5 NOX Value | 528 | 275 | 201 | 98.5%
S8 NOX Value | 531 | 250 | 119 | 96.3%
S12 NOX Value | 332 | 155 | 95 | 95.3%
Chart: 
Chart: 
Sensor 12 NOX values
Sensor 12 NOX values

## Slide 7
Chart: 
Chart: 
O3_2018-06-04-14-32_C
O3_2018-06-08-11-00_C1000ppb
Chart: 
O3_2018-06-11-10-10_Cnofan
Chart: 
O3_2018-06-08-14-36_nofan1000ppb
Sensor 12 NOX values

## Slide 8: O3_2018-06-12-13-30_C
87 | Peak 1 | Peak 2 | Peak 3 | R2
Monitor Concentration | 128.7 ppb | 143.8 ppb | 157.4 ppb | 
S5 NOX Value |  |  |  | ?
S8 NOX Value |  |  |  | ?
S12 NOX Value | 1631 | 1616 | 1590 | ?
Chart: 
AirU sensors were placed in chamber and warmed up. Chamber was flooded with ozone for 50 min prior to running test.